NITP 6.6
National Instrument Test Procedures for 
Automatic Rail Weighbridges

© Commonwealth of Australia 2011

NMI V 6

First edition	—	June 2006

First edition, first revision	—	July 2007

NITP 6.6

First edition	—	December 2011

National Measurement Institute

Bradfield Road, Lindfield, NSW 2070

PO Box 264, Lindfield, NSW 2070

T	(61 2) 8467 3600

F	(61 2) 8467 3610

W	www.measurement.gov.au

Preface

On 30 June 2010 the uniform test procedures (i.e. relevant NMI V documents) were deemed to be national instrument test procedures (NITPs) for the purposes of section 18GG of the National Measurement Act 1960 (Cth).

In 2011 the NITPs were renumbered to better align the numbers with the classes of pattern approval and servicing licensee. As a result this document (NMI V 6) became NITP 6.6.

The only changes that have been made to the latest edition of this document are it has been rebranded, renumbered, renamed and its cross-references have been updated. In all other respects it is identical with NMI V 6.

NMI’s Chief Metrologist has determined that NITP 6.6 contains the test procedures for the verification of automatic rail weighbridges.

Abbreviations

d	scale interval (used to verify/certify a weighbridge)

e	verification scale interval (used to test a control instrument)

E	error (E = P – L)

I	indication

L	load

Lsub	actual calculated value of the substitution load

L	additional load to next changeover point

Max	maximum capacity

Min	minimum capacity

MPE	maximum permissible error

n	total number of wagons in test train

nmax	maximum number of wagons per train

nmin	minimum number of wagons per train

P	indication prior to rounding (P = I + 0.5d – L)

vmax	maximum operating speed

vmin	minimum operating speed

<	less than

	less than or equal to

CONTENTS

Abbreviations	ii

Explanation of Terms	iv

1.	Scope	1

2.	Equipment	1

3.	Visual Inspection	1

3.1	Required Data	1

3.2	Characteristics of the Instrument	2

4.	Test Procedures	2

4.1	Control Instrument	3

4.2	Reference Wagons	4

4.3	Test Train	5

4.4	Weighing Performance	5

4.5	Over-speed Test	7

4.6	Under-speed Test	7

4.7	Roll-back Safeguard Test	7

4.8	Over-weight Wagon Test	7

5.	Suggested Sequence for Testing	7

Appendix A. Test Reports	8

Appendix B. Checklist for Preparation and Planning	18

Appendix C. Worked Examples	19

Explanation of Terms

For explanations of other terms see General Information for Test Procedures.

Adjustment

Alteration of the measurement parameters to bring the instrument within the allowable MPEs for an instrument in use.

Calibration

The set of operations that (under specified conditions) establishes the relationship between the indicated or nominal value of an instrument and the corresponding known value of the measured quantity.

Certification

- The examination of an instrument by a certifier (the holder, or an employee of the holder, of a servicing licence) in order to mark the instrument indicating that it conforms with the relevant test procedures.

- Initial certification is the certification of a new instrument by a certifier, which does not bear a verification or certification mark and has never been verified or certified before.

- Subsequent certification is any certification of an instrument by a certifier because the mark is no longer valid due to such reasons as:

- repairs or adjustments have been made that affect metrological performance; or

- the mark has been defaced or removed.

In-service Inspection

The examination of an instrument by an inspector or certifier to check that:

- the verification or certification mark is valid; and

- the errors do not exceed the MPEs permitted for in-service inspection.

In-service inspection does not permit the instrument to be marked with a verification or certification mark.

Verification

- The examination of an instrument by an inspector in order to mark the instrument indicating that it conforms with the relevant test procedures.

- Initial verification is the verification of a new instrument by an inspector, which does not bear a verification or certification mark and has never been verified or certified before.

- Subsequent verification is any verification of an instrument by an inspector because the mark is no longer valid due to such reasons as:

- repairs or adjustments have been made that affect metrological performance; or

- the mark has been defaced or removed.

- Re-verification is the examination of an instrument by an inspector to check that:

- the verification or certification mark is valid; and

- the instrument has not been modified in any way since verification or certification;

in order to mark the instrument indicating that it conforms with the relevant test procedures.

# Scope

NITP 6.6 describes the test procedures for the verification, certification and in-service inspection of automatic rail weighbridges to ensure that they measure to within the maximum permissible errors specified in the National Measurement Regulations and that they comply with the certificate/s of approval.

These test procedures supersede the dynamic test in Test Procedure 20. Weighing-in-motion Weighing Instruments found in Inspectors Handbook Number 3.

Certificates of approval are based on 
NMI R 106. Automatic Rail Weighbridges. Refer to NMI R 106 for all metrological and technical requirements.

All weighbridges must also comply with the relevant Trade Measurement Act and Regulations.

Refer to NITP 6.1 to 6.4 National Instrument Test Procedures for Non-automatic Weighing Instruments when testing the control instrument.

# Equipment

- Certificate/s of approval.

- Control instrument (see clause 4.1).

- Reference wagons (see clause 4.2).

- Test train (see clause 4.3).

- Test reports (see Appendix A).

# Visual Inspection

Visually inspect the weighbridge and record the following details on Test Report 1.

## Required Data

- Test report reference number.

- Date of the test.

- Type of test: verification, certification or in-service inspection (for in-service inspection ensure that the verification/
certification mark is in place).

- Name of owner/user.

- Address of owner/user.

- Name of contact on site.

- Address where weighbridge is located.

- Manufacturer’s name or mark.

- Importer’s name or mark (if applicable).

- Model.

- Serial number.

- Certificate of approval number.

- Accuracy class for train weighing.

- Accuracy class for wagon weighing.

- Maximum capacity.

- Minimum capacity.

- Scale interval.

- Maximum wagon weight.

- Minimum wagon weight.

- Maximum operating speed.

- Minimum operating speed.

- Maximum number of wagons per train.

- Minimum number of wagons per train.

- Electric power supply voltage.

- Electric power supply frequency.

- Not to be used for liquid products 
(if applicable).

- Direction of weighing (if applicable).

- Scale interval of stationary load 
(if applicable).

- Supplementary and other markings 
(if applicable).

## Characteristics of the Instrument

- Does the weighbridge comply with its certificate/s of approval?

- Is the weighbridge being used in an appropriate manner?

- Are all mandatory descriptive markings clearly and permanently marked on the data plate?

- Is the data plate fixed on the weighing system?

- Is the weighbridge in a suitable operational condition?

- Are there any apparent obstructions to the operation of the weighbridge?

- Is the indicating device accessible during normal operation?

- Is the headwork protected against any other influence likely to affect its performance?

- For rail lines in continuous systems:

- Are the approaches and departures: securely fixed to the railway sleepers; surveyed to confirm the vertical deflection is not excessive; and tamped if survey shows vertical deflection to be excessive?

Note:	Survey points should be on the weighbridge approach, over the weighbridge and also on the weighbridge departures.

- Are the rail sleeper spacings positioned to specifications in relation to the transducers?

- Is there at least 20 mm clearance between the transducer rail and ballast?

- For rail lines in non-continuous systems:

- Are the approaches and departures in the same plane as the weighbridge?

- Are the rail lines fixed securely to the railway sleepers?

- Are the rail lines positioned such that there is no excessive space between the live and dead rails?

- For additional indicating devices: do they exactly repeat the information on the primary indication and does any device for price computation and/or ticket/label printing comply with the requirements of General Supplementary Certificate S1/0/A (or General Supplementary Certificates S1/0 for devices initially verified or certified prior to March 1992)?

# Test Procedures

See Appendix B for a checklist to assist with preparation and planning.

Before testing begins you must:

- discuss the test procedures with the administering authority;

- ensure that adequate skilled personnel are available;

- establish appropriate communication procedures between personnel;

- ensure that all equipment is available and operational; and

- ensure that enough time has been allowed for the testing.

During testing you should:

- avoid any unnecessary commitment of resources; and

- restrict unauthorised access.

The following series of test procedures determine if the performance of a weighbridge meets requirements and whether the weighbridge requires adjustment or service. When adjustment or service is required the testing procedure must be recommenced.

The sequence for testing (also see 
clause 5) is:

- determine if the control instrument is suitable for determining the mass of the reference wagons, and then test the control instrument for compliance with NITP 6.1 to 6.4 (see clause 4.1);

- reference wagons are prepared and the control instrument is used to determine the mass of the reference wagons 
(see clause 4.2);

- the test train is assembled 
(see clause 4.3);

- the weighing performance of the weighbridge is tested in all modes and the results are compared with the mass of the reference wagons when weighed on the control instrument
(see clause 4.4); and

- the weighbridge is tested for over-speed, under-speed, roll-back safeguard and over-weight wagon (see clauses 
4.5 to 4.8).

## Control Instrument

A control instrument is used to determine the mass of the reference wagons which will be used to test the weighbridge.

The control instrument shall:

- be a non-automatic weighing instrument;

- be pattern approved, or have an approved load cell and indicator that comply with General Certificate 6B/0, or have a letter of approval from NMI;

- be capable of having standard masses deposited directly or indirectly on the load receptor easily and in total safety; and

- be capable of weighing the reference wagons uncoupled.

An instrument constructed only for partial weighing may be used as the control instrument as long as end-and-end weighing (see clause 4.2.2) is successfully applied.

Note:	The weighbridge under test may be used as a control instrument as long as it complies with all the requirements for a control instrument.

First, use one of the two methods in Test Report 2 to determine if the control instrument is suitable, and meets the accuracy requirements required for it to determine the mass of the maximum and minimum reference wagons (see 
Appendix C.1 for a worked example). The control instrument shall be at least three or five times more accurate than the weighbridge, namely:

- When the weighbridge under test is used as the control instrument and the weighing of the reference wagons is commenced within 24 hours of testing the control instrument, the control instrument shall be at least three times more accurate than the weighbridge.

- In all other circumstances the control instrument shall be at least at least five times more accurate than the weighbridge. However if the reference wagons are not weighed within 28 days of testing the control instrument you must consult the administering authority. It may be acceptable to use the control instrument if the tests for repeatability and eccentricity are repeated. These are the minimum tests required to provide confidence that the control instrument is still within the required MPE.

Then test the control instrument for compliance with NITP 6.1 to 6.4 National Instrument Test Procedures for Non-automatic Weighing Instruments. The control instrument may be tested:

- to suitable capacities, namely the capacities of the reference wagons determined by clause 4.2.1; or

- to full capacity allowing it to be verified or certified.

Record your results on Test Report 2.

## Reference Wagons

Reference wagons are wagons of known weight used for testing the weighbridge.

Prepare the reference wagons as described in clause 4.2.1 and use the control instrument to determine the mass of the reference wagons (see clause 4.2.2). Record your results on Test Report 1.

### 4.2.1	Preparing the Reference Wagons

A wagon suitable for use as a reference wagon must be prepared in such a manner to prevent loss or gain of material.

Reference wagons shall reflect those normally weighed at the weighbridge, they must not be less than 50 scale intervals, and must include maximum and minimum and cover the range between maximum and minimum wagon weights.

It is acceptable for:

- the maximum reference wagon weight to be up to 10% less than that shown on the data plate; and

- the minimum reference wagon weight to be up to 10% more than that shown on the data plate.

If the maximum and minimum reference wagon weights vary by more than 10% from that shown on the data plate, then the data plate must be adjusted to reflect the actual maximum and minimum reference wagon weights used to make up the test train.

Reference wagons must be protected during storage and transportation in order to ensure that there is no introduction of errors, from external factors, e.g. rain, from the time it becomes a reference wagon until the time it is used to test the weighbridge.

Check with the administering authority to determine the period of validity of a reference wagon. This will be based on the material used, the possibility of the introduction of errors due to influence factors and the time taken to complete testing. Where reference wagons are covered by a Regulation 13 certificate, the certificate will indicate the time interval.

In some rail systems a type of compound wagon is used but these are generally not used as a reference wagon. Figure 1 shows a typical compound wagon with three containers also referred to as a ‘multi-pack wagon’.

Figure 1. A typical compound wagon with three containers

As bogies within a typical compound wagon share the load from the adjacent containers, it is not possible to assign a weight to each individual container and the compound wagon must be treated as a single reference wagon.

### 4.2.2	Determining the Mass of the Reference Wagons

The preferred method of determining the mass of the reference wagons is by direct weighing of uncoupled wagons, although end-on-end weighing may be used as long as you consult the administering authority.

End-and-end weighing applies to instruments that operate by partial weighing of wagons. When using end-and-end weighing ensure that:

- the wagon fits within the apron and load receptor;

- both rails along the length of the apron and load receptor are in the same plane;

- the top surface of both rails along the length of the weigh zone are vertical and level aligned to ±1 mm; and

- the alignment has been checked along both rails at not less than two positions on the load receptor and not less than two positions within a wagon length from the load receptor on each associated apron.

See Appendix C.2 for a worked example and instructions on how to conduct end-and-end weighing.

## Test Train

Assemble a test train that reflects a typical train weighed by the weighbridge, and record the makeup of the train on Test Report 1.

The test train may consist of: locomotive/s, non-reference wagons and reference wagons, or locomotive/s and reference wagons.

More than one test train may be required, 
e.g. where a typical train is weighed in one direction full and in the other direction empty.

Wagons in a typical train may have a:

- full or nearly full load; or

- mixed load which covers the range from minimum to maximum load as long as no interval between load bands is greater than 20% of the maximum wagon weight; or

- empty or relatively empty load.

If the weighbridge is used as a multipurpose weighbridge a mixed train can be used for testing.

The number of reference wagons in a test train shall be in accordance with Table 1.

Table 1. Number of reference wagons required

| Total number of wagons in test train (n) | Minimum number of reference wagons |
| --- | --- |
| < n  10 | n |
| 10 < n  30 | 10 |
| 30 < n | 15 |

If the number of reference wagons is less than the total number of wagons in a test train, the reference wagons shall be distributed evenly throughout the train in the following positions:

- after the locomotive(s) and the first two wagons;

- after 25% of the wagons;

- after 50% of the wagons;

- after 75% of the wagons; and

- at the end of the train.

Where the weight of the reference wagons is variable, there shall be a proportional representation of reference wagons uniformly distributed through the test train.

## Weighing Performance

The purpose of this test is to record the weighing performance of the weighbridge at near-maximum and near-minimum operating speeds and in all modes, and to compare the results with the mass of the reference wagons when weighed on the control instrument.

When the typical train operates in more than one mode (direction, speed, loading) the weighbridge shall be tested in each mode, namely if it:

- is used in both directions it shall be tested for each directional mode;

- operates over a range of speeds it shall be tested in each speed band, and an operational speed band shall not exceed 20 km/h, e.g. when the maximum operating speed is 80 km/h and the minimum operating speed is 20 km/h, then a separate weighing performance test shall be carried out for each speed band 20 to 40 km/h, 40 to 60 km/h, and
60 to 80 km/h); and

- operates with variable loads then it shall be tested in each loading mode, e.g. all full, all empty, mixed, partially loaded.

Notes:

Ensure that the weight of any vehicle not intended to be weighed (e.g. locomotives) is not recorded.

Where the weighbridge has facilities for calculating net mass from gross wagon weights and the tare mass of the wagons, ensure that these calculations are correct.

If site conditions limit the ability to achieve the operating speeds, capacity and wagon weights approved in the certificate/s of approval, those speeds, capacities and weights specific to the site shall be marked on the data plate.

To avoid duplication, the results of tests performed for pattern approval may be used for initial verification/certification as long as you consult the administering authority and NMI.

- Run the test train over the weighbridge at a chosen speed (see clause 4.4).

- Record the indicated weight of each reference wagon as displayed or printed by the weighbridge, on Test Report 1.

- Repeat steps 1 to 2 until a minimum of three runs and a minimum of 60 reference wagon weights have been recorded. Note that if a mixed train is used, a minimum of 12 reference wagon weighs shall be recorded for each load band within the range.

- Use the results from step 3 to calculate the errors for weighing the reference wagons and the test train (see clauses 4.4.1 and 4.4.2) respectively. Record your results on Test Report 1.

Note:	The particular system under test may calculate the errors automatically.

- Repeat steps 1 to 4 for all modes.

### 4.4.1	Wagon Weighing

To calculate the error for each reference wagon:

(i)	Calculate the allowable MPE for wagon weighing. It shall be the greatest of the following values:

- the value calculated according to 
Table 2, rounded to the nearest scale interval; or

- the value calculated according to 
Table 2, rounded to the nearest scale interval for the weight of a single wagon equal to 35% of the maximum wagon weight (as marked on the data plate); or

-  one scale interval.

See Appendixes C.3.1(a) and C.3.2(a) for two worked examples.

          - Calculate the actual error by subtracting the indicated weight from the weight recorded by the control instrument.

          - For each run, calculate the percentage of reference wagons within both MPE and twice MPE. For verification and certification, it is acceptable for up to 10% of the weighing results taken from one or more passes of the test train to exceed the appropriate MPE given in Table 2 but they shall not exceed two times that value.

    - Record all your results on Test Report 1.

Table 2. MPEs for weighbridges

| Accuracy class | Percentage of mass of single reference wagon or test train |  |
| --- | --- | --- |
|  | Verification/ certification | In-service inspection |
| 0.2 | ±0.10% | ±0.2% |
| 0.5 | ±0.25% | ±0.5% |
| 1 | ±0.50% | ±1.0% |
| 2 | ±1.00% | ±2.0% |

### 4.4.2	Train Weighing

To calculate the error for the test train:

(i)	Calculate the allowable MPE for train weighing. It shall be the greatest of the following values:

- the value calculated according to 
Table 2, rounded to the nearest scale interval; or

- the value calculated according to 
Table 2 of the weight of a single wagon equal to 35% of the maximum wagon weight (as marked on the data plate) multiplied by the number of reference wagons in the train (not exceeding 
10 wagons) and rounded to the nearest scale interval; or

-  one scale interval for each wagon in the train, but not exceeding ten scale intervals.

See Appendixes C.3.1(b) and C.3.2(b) for two worked examples.

          - Total the weight of the reference wagons recorded by the control instrument.

          - Total the indicated weight of all the reference wagons in each pass of the test train.

          - Subtract the value obtained in (iii) from the value obtained in (ii).

          - Ensure the error is within the MPE.

          - Record all your results on Test Report 1.

## Over-speed Test

The purpose of this test, which may be simulated and only needs to be run in one mode, is to check that the over-speed protection system inhibits printing a weight when the maximum speed for which the weighbridge is approved, is exceeded.

- Run the test train over the weighbridge at a speed just over (approximately 10%) maximum operating speed.

- Check that the over-speed protection system inhibits printing a weight when the maximum speed is exceeded.

- Record your results on Test Report 1.

## Under-speed Test

The purpose of this test, which only needs to be run in one mode, is to check that the under-speed protection system inhibits printing a weight when the minimum speed for which the weighbridge is approved is not achieved.

- Run the test train over the weighbridge at a speed just below (approximately 10%) minimum operating speed.

- Check that the under-speed protection system inhibits printing a weight when the minimum operating speed is not achieved.

- Record your results on Test Report 1.

## Roll-back Safeguard Test

The purpose of this test, which only needs to be run in one mode, is to check that the roll-back safeguard is detected automatically and that the weighing process is either aborted or continued from the point reached before roll-back commenced.

- Run approximately 75% of the test train over the weighbridge and stop the test train.

- Reverse the test train for approximately 25% of its length.

- Recommence the test by running the test train forward.

- Check that the roll-back safeguard is detected automatically and that the weighing process is either aborted or continued from the point reached before roll-back commenced.

- Record your results on Test Report 1.

## Over-weight Wagon Test

The purpose of this test is to check that the printer does not record an overweight wagon.

In view of the occupational health and safety ramifications, this test should only be conducted with the authorisation of the rail operator. An acceptable solution is to simulate this test.

- Load any wagon to a capacity greater than the maximum capacity plus nine scale intervals.

- Include this wagon in the test train.

- Conduct a test run.

- Check that the weight of the over-weight wagon has not been recorded.

- Record your results on Test Report 1.

# Suggested Sequence for Testing

- Before testing begins:

- discuss the test procedures with the administering authority;

- ensure that adequate skilled personnel are available;

- establish appropriate communication procedures between personnel;

- ensure that all equipment is available and operational; and

- ensure that enough time has been allowed for the testing.

Remember to avoid any unnecessary commitment of resources, and to restrict unauthorised access.

- Check the certificate/s of approval for any additional tests required. Make provision for including these tests in the testing sequence.

- Visually inspect the weighbridge and record the required details on Test Report 1.

- Determine if the control instrument is suitable, and meets the accuracy requirements required for it to determine the mass of the reference wagons (see clause 4.1). Record the results on Test Report 2.

- Test the control instrument for compliance with NITP 6.1 to 6.4 
(see clause 4.1) and record the results on Test Report 2.

- Prepare the reference wagons 
(see clause 4.2).

- Use the control instrument to determine the mass of the reference wagons 
(see clause 4.2) and record the results on Test Report 1.

- Assemble the test train (see clause 4.3) and record the makeup of the train on Test Report 1.

- Test the weighing performance of the weighbridge in all modes (see clause 4.4) and record the results (for a minimum of 60 reference wagons over three runs) on Test Report 1.

- Calculate the allowable MPEs for wagon weighing (see clause 4.4.1(i)) and train weighing (see clause 4.4.2(i)) and record the results on Test Report 1.

- Calculate the actual errors for wagon weighing (see clause 4.4.1) and train weighing (see clause 4.4.2) and record the results on Test Report 1.

- Compare the allowable MPEs for wagon weighing and train weighing with the actual errors and record the results on Test Report 1.

- Test the weighbridge to see if it records over-speed (see clause 4.5) and record the results on Test Report 1.

- Test the weighbridge to see if it records under-speed (see clause 4.6) and record the results on Test Report 1.

- Test the weighbridge to see if it detects roll-back (see clauses 4.7) and record the results on Test Report 1.

- Test the weighbridge to see if it records an over-weight wagon (see clause 4.8) and record the results on Test Report 1.

- Determine whether the weighbridge has passed or failed.

Note:	If the weighbridge fails one of the modes it may still be stamped providing it has passed other modes and the data plate is changed to reflect this.

- Complete Test Report 1 and Test Report 2.

- Carry out anything else you need to do to complete the procedure. This may include:

- obliterating verification, certification and control marks from the weighbridge; and

- stamping the weighbridge (for more information on stamping see
General Information for Test Procedures).

Appendix A. Test Reports

Appendix A contains two test reports:

- Test Report 1 is for weighbridges; and

- Test Report 2 is for control instruments.

Although the format of the test reports may vary according to the individual needs and requirements of trade measurement authorities and licensees, the following test reports contains the minimum amount of information that must be recorded.

If the certificate/s of approval requires additional tests, attach pages that record the results of these tests.

Number each page of the test report in the style shown at the top of each page.

Test report reference number	Date of test

Type of test (tick one)	 Verification	 Certification	 In-service inspection

For in-service inspection record the verification/certification mark

Name of owner/user

Address of owner/user

Name of contact on site

Address where weighbridge is located

Manufacturer’s name or mark

Importer’s name or mark (if applicable)

Model		Serial number

Certificate/s of approval number

Accuracy class for wagon weighing (circle correct one)	0.2	0.5	1	2

Accuracy class for train weighing (circle correct one)	0.2	0.5	1	2

Accuracy class, other (if applicable)

Maximum capacity (Max)	kg or t	Minimum capacity (Min)	kg or t

Scale interval (d)	kg or t

Maximum wagon weight	kg or t

Minimum wagon weight.	kg or t

Maximum operating speed (vmax)	km/h

Minimum operating speed (vmin)	km/h

Maximum number of wagons per train (nmax)

Minimum number of wagons per train (nmin)

Electric power supply voltage	V	Electric power supply frequency	Hz

Not to be used for liquid products (if applicable)

Direction of weighing (if applicable)

Scale interval of stationary load (if applicable)	kg or t

Supplementary and other markings (if applicable)

| Does the weighbridge comply with its certificate/s of approval? | yes/no |
| --- | --- |
| Is the weighbridge used in accordance with its certificate/s of approval? | yes/no |
| Are all mandatory descriptive markings clearly and permanently marked on data plate? | yes/no |
| Is the weighbridge in a suitable operational condition? | yes/no |
| Are there any apparent obstructions to the operation of the weighbridge? | yes/no |
| Is the indicating device accessible during normal operation? | yes/no |
| Is the headwork protected against any other influence likely to affect its performance? | yes/no |
| For the rail line in continuous systems — are the approaches and departures: securely fixed to the railway sleepers, surveyed regularly to confirm the vertical deflection is not excessive, and tamped if survey shows vertical deflection to be excessive? | yes/no/na |
| For the rail line in continuous systems — are the rail sleeper spacings positioned to specifications in relation to the transducers? | yes/no/na |
| For the rail line in continuous systems — is there at least 20 mm clearance between the transducer rail and ballast? | yes/no/na |
| For the rail line in non-continuous systems — are the approaches and departures in the same plane as the weighbridge? | yes/no/na |
| For the rail line in non-continuous systems — are the approaches and departures in the same plane as the weighbridge? | yes/no/na |
| For the rail line in non-continuous systems — are the rail lines positioned such that there is no excessive space between the live and dead rails? | yes/no/na |
| Are the rail sleeper spacings positioned to specifications in relation to the transducers? | yes/no/na |
| Is there at least 20 mm clearance between the transducer rail and ballast? | yes/no/na |
| For additional indicating devices: do they exactly repeat the information on the primary indication and does any device for price computation and/or ticket/label printing comply with the requirements of the General Supplementary Certificates S1/0/A or S1/0? | yes/no/na |

Weighing Performance

Number of reference wagons	 Number of non-reference wagons

Number of locomotives	 Position of locomotives

Number and description of directions

Maximum operating speed	km/h	Minimum operating speed	 km/h

Other speeds: 	 km/h   	 km/h  	 km/h   	 km/h

Full load	 t or kg	 Empty load	 t or kg

Other loads	 t or kg   	 t or kg   	 t or kg   	 t or kg

Allowable MPE for wagon weighing
(greatest value below, rounded to nearest scale interval)    ±	 t or kg

| MPE from Table 2  reference wagon weight | Reference wagon 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 Sum of all ref wagons | ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg ± t or kg |
| --- | --- | --- |
| MPE from Table 2  35% of maximum wagon weight marked on the data plate |  | ± t or kg |
|  one scale interval |  | ± t or kg |

Allowable MPE for train weighing
(greatest value below, rounded to nearest scale interval)    ±	 t or kg

| MPE from Table 2  sum of all the reference wagon weights | ± t or kg |
| --- | --- |
| MPE from Table 2  35% of maximum wagon weight marked on the data plate  number of reference wagons (but not exceeding ten) | ± t or kg |
|  one scale interval, but not exceeding ten scale intervals | ± t or kg |

Copy this page for each mode being tested

Direction	 	Speed	km/h

Reference wagon load	 empty load	 full load 	 mixed load

| Reference wagon | MPE | ID | Position in train | Using the control instrument | Run 1 (t or kg) |  | Run 2 (t or kg) |  | Run 3 (t or kg) |  | Run 4 (t or kg) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Weight (t or kg) | Indicated weight | Error | Indicated weight | Error | Indicated weight | Error | Indicated weight | Error |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |  |  |  |  |  |
| 16 |  |  |  |  |  |  |  |  |  |  |  |  |
| 17 |  |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  |  |  |  |  |  |  |  |  |  |  |  |
| 19 |  |  |  |  |  |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |  |  |  |  |  |
| Total train weight |  |  |  |  |  |  |  |  |  |  |  |  |

Wagon weighing	Allowable MPE (from previous page) ±	t or k

Percentage of reference wagons within MPE	%

Percentage of reference wagons within twice MPE	% (up to 10% acceptable)

Train weighing	Allowable MPE (from previous page) ±	t or k

Test Results

| Weighing performance — wagon weighing (clause 4.4.1) |  Pass  Fail |
| --- | --- |
| Weighing performance — train weighing (clause 4.4.2) |  Pass  Fail |
| Over-speed test (clause 4.5) |  Pass  Fail |
| Under-speed test (clause 4.6) |  Pass  Fail |
| Roll-back test (clause 4.7) |  Pass  Fail |
| Over-weight test (clause 4.8) |  Pass  Fail |
| Overall result |  Pass  Fail |

Inspector’s/certifier’s name	Identification number

Signature

Comments

Calculation to Determine if the Control Instrument is Suitable for Use
and Meets the Accuracy Requirements

| Proposed reference wagon weights |  |  |
| --- | --- | --- |
| Minimum capacity (kg or t) | Intermediate capacities as applicable (kg or t) | Maximum capacity (kg or t) |

Method 1 — Calculation NOT using Change Points

| Part 1. Accuracy requirements for the control instrument [(MPE% / 100)  St  ⅓] |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Ref wagon ID | Mass of reference wagon | Accuracy class for wagon weighing |  | Absolute MPE from Table 2 | MPE of reference wagon | Accuracy requirement (1/3 MPE of reference wagon) |
| 1 |  |  |  | ± | ± | ± |
| 2 |  |  |  | ± | ± | ± |
| 3 |  |  |  | ± | ± | ± |
| Part 2. MPE of the control instrument |  |  |  |  |  |  |
| Ref wagon ID | Mass of reference wagon | Verification scale interval (e) | Number of verification scale intervals |  | MPE from Table 3 | MPE value |
| 1 |  |  |  |  | ± | ± |
| 2 |  |  |  |  | ± | ± |
| 3 |  |  |  |  | ± | ± |
| Does the control instrument meet the accuracy requirements? Note: MPE of control instrument must be ⅓ MPE of reference wagon |  |  |  |  |  | yes/no |

Table 3. MPEs for loads, m, expressed in verification scale intervals, e

| MPEs | Class 3 | Class 4 |
| --- | --- | --- |
| ±0.5e | 0 ≤ m ≤ 500 | 0 ≤ m ≤ 50 |
| ±1e | 500 < m ≤ 2 000 | 50 < m ≤ 200 |
| ±1.5e | 2 00 < m ≤ 10 000 | 200 < m ≤ 1 000 |

Method 2 — Calculation using Change Points

| Ref wagon ID | P = I + 0.5d – L |  |  | E = P – L |  |  | % of L (t or kg) | Maximum error acceptable (⅓ or 1/5, circle which applies) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | I | 0.5d | ΔL | P | L(=I) | E |  |  |
| 1 |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |
| Does the control instrument meet the accuracy requirements? |  |  |  |  |  |  | yes/no |  |

Refer to NITP 6.1 to 6.4 for the test procedure

Test report reference number	Date of test

Type of test (tick one)	 Verification	 Certification

 Other (attach either a 6B/0 analysis or a letter of approval from NMI)

Name of owner/user

Address of owner/user

Name of contact on premises

Address where instrument located

Description of instrument

Manufacturer/s	Model	Instrument serial number

Certificate/s of approval number (if applicable)

Maximum capacity (Max)	Minimum capacity (Min)

Verification scale interval (e)	Accuracy class

| Does the instrument comply with its certificate/s of approval? | yes/no/na |
| --- | --- |
| Is the instrument being used in an appropriate manner? | yes/no |
| Are all mandatory descriptive markings clearly and permanently marked on the data plate? | yes/no |
| Is the data plate fixed on the instrument? | yes/no |
| Is the instrument complete? | yes/no |
| Is the instrument broken? | yes/no |
| Is the instrument clean? | yes/no |
| Is the instrument operational? | yes/no |
| Are there any apparent obstructions to the operation of the instrument? | yes/no |
| Does the operator (and where applicable, the customer) have a clear and unobstructed view of the indicating device and the whole weighing operation? | yes/no |
| Is the instrument adequately protected against abnormal dust, air movement, vibrations, atmospheric conditions and any other influence likely to affect its performance? | yes/no |
| If applicable, does the steelyard, tare bar or proportional weight comply with the mandatory requirements in respect to design and marking? | yes/no/na |
| Does the weighbridge comply with the relevant Trade Measurement (Weighbridge) Regulations and the Code of Practice for Weighbridge Installations? | yes/no |
| For additional indicating devices: do they exactly repeat the information on the primary indication and does any device for price computation and/or ticket/label printing comply with the requirements of General Supplementary Certificate S1/0/A (or General Supplementary Certificates S1/0 and S2/0 for devices initially verified or certified prior to March 1992)? | yes/no/na |

Test Results

| Repeatability (NITP 6.1 to 6.4, clause 5.1) | Load |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | First reading |  |  |  |  |
|  | Second reading |  |  |  |  |
|  | Third reading |  |  |  |  |
|  | Difference |  |  |  |  |
|  |  Pass  Fail |  |  |  |  |
| Eccentricity (NITP 6.1 to 6.4, clause 5.2) | Number of supports: |  |  |  |  |
|  | Load used: |  |  |  |  |
|  | Position 1 |  |  | Position 7 |  |
|  | Position 2 |  |  | Position 8 |  |
|  | Position 3 |  |  | Position 9 |  |
|  | Position 4 |  |  | Position 10 |  |
|  | Position 5 |  |  | Position 11 |  |
|  | Position 6 |  |  | Position 12 |  |
|  |  Pass  Fail |  |  |  |  |
| Zero setting (NITP 6.1 to 6.4, clause 5.3) |  Pass  Fail |  |  |  |  |
| Weighing performance not using substitution load (NITP 6.1 to 6.4, clause 5.4.1) Note: For weighing performance using substitution load refer to the next page | Loads applied (minimum 5) |  |  | Indication up | Indication down |
|  | Over-range blanking  Pass  Fail  na |  |  |  |  |
|  |  Pass  Fail  na |  |  |  |  |

| Weighing performance using substitution load (NITP 6.1 to 6.4, clause 5.4.2) |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Method used  Method A  Method B |  |  |  |  |  |  |  |  |  |  |
| MPE change points |  |  |  |  |  |  |  |  |  |  |
| Available standard weights |  |  |  |  |  |  |  |  |  |  |
| First substitution load |  |  |  |  |  |  |  |  |  |  |
| Second substitution load |  |  |  |  |  |  |  |  |  |  |
| Third substitution load |  |  |  |  |  |  |  |  |  |  |
| Up | L | Makeup of load | MPE | I | ½e | ΔL | E | Lsub | Lsub (rounded) | Pass/fail/na |
| Over-range blanking  Pass  Fail  na |  |  |  |  |  |  |  |  |  |  |
| Down | L |  | Makeup of load |  | MPE |  | I |  | Pass/fail |  |

| Discrimination (NITP 6.1 to 6.4, clause 5.5) |  Pass  Fail |
| --- | --- |
| Sensitivity (NITP 6.1 to 6.4, clause 5.6) |  Pass  Fail  na |
| Accuracy of tare setting (NITP 6.1 to 6.4, clause 5.7) |  Pass  Fail  na |
| Price computation (NITP 6.1 to 6.4, clause 5.8) |  Pass  Fail  na |
| Overall result |  Pass  Fail |

Inspector’s/certifier’s name	Identification number

Signature

Comments

Appendix B. Checklist for Preparation and Planning

| Does the weighbridge comply with its certificate of approval? | yes/no |
| --- | --- |
| Does the certificate of approval contain any special conditions? | yes/no |
| Does the certificate of approval contain any additional test/s? | yes/no |
| Is written confirmation from NMI required for the control instrument? | yes/no |
| Has the relevant administering authority been advised? | yes/no |
| Have arrangements been made with the site operator to ensure exclusive availability of the weighbridge during testing? | yes/no |
| Have arrangements been made with the site operator to ensure exclusive availability of the control instrument during testing? | yes/no |
| Is there sufficient availability of reference standards of measurement to test the control instrument to the maximum reference wagons weight? | yes/no |
| Do the reference standards of measurement have a current Regulation 13 certificate for inspectors’ reference standards class 3? | yes/no |
| Have arrangements been made to test the control instrument within the prescribed time? | yes/no |
| Has the suitability of the control instrument been determined? | yes/no |
| Have arrangements been made to ensure the test results for the control instrument are included in the test report for the automatic rail weighbridge? | yes/no |
| Has the number of reference wagons been determined? | yes/no |
| Has the mass of the reference wagons been determined? | yes/no |
| Has the number of runs required been determined? | yes/no |
| Has the configuration of the test train been arranged? | yes/no |
| Has the MPE for wagon weighing been calculated? | yes/no |
| Has the MPE for train weighing been calculated? | yes/no |
| Have the required number of skilled personnel been scheduled to assist? | yes/no |
| Are suitable communication systems in place? | yes/no |
| Is all the required equipment operational and available? | yes/no |

Appendix C. Worked Examples

C.1	How to Determine the Suitability of the Control Instrument

In this example the weighbridge has maximum wagon weight of 80 000 kg and a minimum wagon weight of 18 000 kg. The weighbridge is also capable of being used as a non-automatic weighbridge.

The non-automatic weighbridge has a capacity of 100 000 t with a verification scale interval of 50 kg. It will be tested within 24 hours and can be used as the control instrument provided it meets the accuracy requirements. The accuracy class of the control instrument is class 3.

The reference wagon weights used are
18 000 kg and 78 000 kg. The accuracy class for wagon weighing of the weighbridge is 0.5 (MPE of ±0.25%).

In this case, the maximum error acceptable for the control instrument must be equal to or less than one-third the MPE value of the reference wagon. This accuracy only applies when the weighbridge is used as the control instrument and the preparation of the reference wagons is commenced within 
24 hours of testing the control instrument.

This example provides two different methods to determine the suitability of the control instrument. If method 1 does not show the control instrument is suitable then method 2 can be employed to ensure the control instrument is suitable.

C.1.1	Method 1 — Calculation NOT using Change Points

If this method determines the control instrument is suitable for both high and low capacity reference wagons it will also be suitable for all reference wagons within the range. The weighing results for each reference wagon are taken directly from the indication of the control instrument without the need to use change points as described in method 2.

Table 4 outlines the calculations required to determine suitability of the control instrument without the need to use change points. In part 1 the weighbridge specifications are used to calculate the maximum error acceptable; and in part 2 the control instrument specifications are used to calculate the MPE mass value for the control instrument. The results from parts 1 and 2 are used to determine if the control instrument is suitable.

Table 4. Calculation NOT using change points

| Part 1. Accuracy requirements for the control instrument [(MPE% / 100)  St  ⅓] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Ref wagon ID | Mass of reference wagon | Accuracy class for wagon weighing | Absolute MPE from Table 2 | MPE of reference wagon | Accuracy requirement (1/3 MPE of reference wagon) |
| 1 | 18 000 kg | 0.5 | ±0.25% | ±45 kg | ±15 kg |
| 2 | 78 000 kg | 0.5 | ±0.25% | ±195 kg | ±65 kg |
| Part 2. MPE of the control instrument |  |  |  |  |  |
| Ref wagon ID | Mass of reference wagon | Verification scale interval | Number of verification scale intervals | MPE from Table 3 | MPE value |
| 1 | 18 000 kg | 50 kg | 360 | ±0.5e | ±25 kg |
| 2 | 78 000 kg | 50 kg | 1 560 | ±1e | ±50 kg |
| Does the control instrument meet the accuracy requirements? |  |  |  |  | No |

The calculations in Table 4 are performed as follows.

In part 1 the mass of the reference wagon is multiplied by the absolute MPE for the weighbridge from Table 2 to give the MPE of the reference wagons. The result is then divided by one-third. This is the maximum error acceptable for the accuracy for the control instrument.

In part 2 the mass of each reference wagon is divided by the verification scale interval to provide the number of verification scale intervals. This is then used to determine the MPE. The MPE is multiplied by the verification scale interval to provide the MPE value for the control instrument at the mass of the reference wagons.

The results in Table 4 are interpreted as follows:

- in part 1 for the 18 000 kg reference wagon the accuracy requirement calculated for the control instrument is ±15 kg;

- in part 2 at 18 000 kg the MPE of the control instrument is ±25 kg.

Therefore the control instrument is unacceptable as the ±25 kg mass value of the control instrument is greater than the accuracy requirement of ±15 kg for that particular reference wagon.

In part 1 for the 78 000 kg reference wagon the accuracy requirement calculated for the control instrument is ±65 kg.

In part 2 at 78 000 kg the MPE of the control instrument is ±50 kg.

Therefore the control instrument is acceptable as the ±50 kg mass value of the control instrument is less than the accuracy requirement of ±65 kg for that particular reference wagon.

As a result of these calculations the control instrument is only suitable for 78 000 kg reference wagons.

This example demonstrates that where there is significance difference in the mass between low capacity reference wagons and high capacity reference wagons a single interval control instrument may not be suitable. However a multi-interval instrument may have proved to be suitable.

If an 18 000 kg reference wagon is required then the use of change points as shown in method 2 will determine suitability.

C.1.2	Method 2 — Calculation using Change Points

Change points are used to obtain greater resolution of the control instrument when the scale interval of the control instrument does not meet the one-third or one-fifth accuracy requirement.

This example demonstrates how change points can be applied using the same data used in method 1. If change point weights are used they must be used to determine all reference wagons and be recorded as part of the test report. Every reference wagon weighing calculation must show the control instrument is operating within the maximum error acceptable.

The error is determined by following the same basic principal used for substitution load as outlined in NITP 6.1 to 6.4 National Instrument Test Procedures for Non-automatic Weighing Instruments.

At a certain load, L, the indicated value, I, is noted.

Additional standard weights of 0.1e are successively added to the load receptor until the indication of the instrument is increased unambiguously by one scale interval (I + e). The total mass of the additional weights is DL.

The true indication, P, prior to rounding is found by using the following formula:

P = I + 0.5e – DL

The error, E, prior to rounding is E = P – L.

Thus E = (I + 0.5e – DL) – L.

For the 18 000 kg reference wagon: after adding successive standard weights of 10 kg, the indication changes up 1e from 18 000 kg to 18 050 kg at an additional load of 10 kg. Inserted in the above formula these observations give:

P = (18 000 + 25 – 10) kg = 18 015 kg

Thus the true indication prior to rounding is 18 015 kg and the error is:

E = (18 015 – 18 000) kg = 15 kg

The MPE from Table 2 is ±0.25%. The allowable error for an 18 000 kg reference wagon is:

±0.25% of 18 000 kg = ±45 kg

The accuracy requirement for the control instrument is ±45 ÷ 3 = ±15 kg. Since the error in the control instrument is known and is less than or equal to 15 kg the control instrument is suitable.

For the 78 000 kg reference wagon: after adding successive standard weights of 10 kg, the indication changes up 1e from 78 000 kg to 78 050 kg at an additional load of 20 kg. Inserted in the above formula these observations give:

P = (78 000 + 25 – 20) kg = 78 005 kg

Thus the true indication prior to rounding is 78 005 kg and the error for the control instrument at the mass of the reference wagon is:

E = (78 005 – 78 000) kg = 5 kg

The MPE from Table 2 is ±0.25%. The allowable error for a 78 000 kg reference wagon is:

±0.25% of 78 000 kg = ±195 kg

The accuracy requirement for the control instrument is ±195 ÷ 3 = ±65 kg. Since the error in the control instrument is known and is less than ±65 kg the control instrument is suitable.

This calculation must be carried out for every wagon weight used in the test. Record results in a table similar to 
Table 5.

C.2	How to Determine the Mass of Reference Wagons using End-and-End Weighing

- Weigh each bogie in the centre and at each end of the load receptor and in both directions.

- Add the six weight indications for direction 1 and divide the result by three.

- Add the six weight indications for direction 2 and divide the result by three.

- Add the results obtained in 2 and 3 and divided by two.

See Table 6 for an example of an end-and-end calibration test sheet.

Table 5. Calculation using change points

| Ref wagon ID | P = I + 0.5d – L |  |  | E = P – L |  |  | % of L (t or kg) | Maximum error acceptable ⅓ |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | I | 0.5d | ΔL | P | L(=I) | E |  |  |
| 1 | 18 000 kg | 25 kg | 10 kg | 18 015 kg | 18 000 | 15 kg | ±45 kg | ±15 kg |
| 2 | 78 000 kg | 25 kg | 20 kg | 78 005 kg | 78 000 | 5 kg | ±195 kg | ±65 kg |
| Does the control instrument meet the accuracy requirements? |  |  |  |  |  |  | Yes |  |

Table 6. Example of an end-and-end calibration test sheet

|  | Position on load receptor | Indicated weight (t) |  |
| --- | --- | --- | --- |
|  |  | Direction 1 | Direction 2 |
| First bogie | Leading end Centre Trailing end | 14.76 14.75 14.75 | 14.27 14.26 14.26 |
| Second bogie | Leading end Centre Trailing end | 14.75 14.75 14.74 | 14.25 14.25 14.24 |
| Total of six weighings |  | 88.50 | 85.53 |
| Divide total by three |  | 29.50 | 28.51 |
| Total for both directions |  | 58.01 |  |
| Divide total by two |  | 29.00 |  |
| Weight of the reference wagon |  | 29.00 |  |

C.3	How to Calculate the MPEs for Weighing Performance

C.3.1	Example 1

In this example the data plate on the weighbridge, which is to be certified, is marked as follows:

- class 1 for train weighing

- class 2 for wagon weighing

- verification scale interval of 0.2 t

- maximum wagon weight of 100 t

The test train consists of 58 loaded wagons in total, 15 of which are reference wagons with a combined reference wagon weight of 1 380 t. The 15 reference wagon weights vary between 80 t and 93.6 t.

(a)	Wagon Weighing

The following calculations are used to determine the allowable MPE for a single reference wagon weighing 93.6 t. It is the greatest of the values calculated below, rounded to the nearest scale interval.

- The value calculated according to Table 2.

Reference wagon wt	=	93.6 t

MPE from Table 2	=	±1.0%

Mass value of MPE	=	±1.0%  93.6 t

=	±0.936 t

- The value calculated according to Table 2, for the weight of a single wagon equal to 35% of the maximum wagon weight (as marked on the data plate).

Maximum wagon wt	=	100 t

35% of max wagon wt	=	35 t

MPE from Table 2	=	±1.0%

Mass value of MPE	=	±1.0%  35 t

=	±0.35 t

- ± One scale interval.

One scale interval	=	0.2 t

Mass value of MPE	=	±0.2 t

Therefore the MPE calculated in point 1, rounded to the nearest scale interval, is the allowable MPE for the 93.6 t reference wagon, i.e. ±1 t. Table 7 also outlines these calculations.

Note:	You have to calculate the allowable MPE for each reference wagon.

At verification/certification, it is acceptable for up to 10% of the weighing results taken from one or more passes of the test train to exceed the appropriate MPE in Table 2 but they shall not exceed two times that value. Therefore:

- 90% of the reference wagons (14) can have a maximum error of ±1 t; and

- 10% of the reference wagons (1) can have a maximum error of ±2 t.

Table 7. Allowable MPE for wagon weighing (Example 1)
(greatest value below, rounded to nearest scale interval) = ±1 t

| MPE from Table 2  reference wagon weight = ±1.0%  93.6 t = | ±0.936 t |
| --- | --- |
| MPE from Table 2  35% of maximum wagon weight marked on the data plate = ±1.0%  100 t  35/100 = | ±0.35 t |
|  one scale interval = | ±0.2 t |

Table 8. Allowable MPE for train weighing (Example 1)
(greatest value below, rounded to nearest scale interval) = ±7 t

| MPE from Table 2  sum of all the reference wagon weights = ±0.5%  1 380 t = | ±6.9 t |
| --- | --- |
| MPE from Table 2  35% of maximum wagon weight marked on the data plate  number of reference wagons (but not exceeding ten) ±0.5%  100 t  35/100  10 = | ±1.75 t |
|  one scale interval, but not exceeding ten scale intervals = ±0.2 t  10 = | ±2 t |

(b)	Train Weighing

The allowable MPE for the total weight of the reference wagons is the greatest of the values calculated below, rounded to the nearest scale interval.

- The value calculated according to Table 2.

Combined ref wgn wt	=	1 380 t

MPE from Table 2	=	±0.5%

Mass value of MPE 	=	±0.5%  1 380 t

=	±6.9 t

- The value calculated according to Table 2 of the weight of a single wagon equal to 35% of the maximum wagon weight (as marked on the data plate) multiplied by the number of reference wagons in the test train (not exceeding 10 wagons).

Maximum wagon wt	=	100 t

35% of max wagon wt	=	35 t

MPE from Table 2	=	±0.5%

Mass value of MPE	=	±0.5%  35 t

=	±0.175 t

Multiplied by not more than 10 wagons
=	±0.175 t 10	=	±1.75 t

- ± One scale interval for each wagon in the test train (but not exceeding 10 scale intervals).

One scale interval	=	0.2 t

Mass value of MPE	=	±0.2 t  10

=	±2 t

Therefore the MPE calculated in point 1, rounded to the nearest scale interval, is the allowable MPE for the total weight of the reference wagons, i.e. ±7 t. Table 8 also outlines these calculations.

C.3.2	Example 2

In this example the data plate on the weighbridge, which is to be certified, is marked as follows:

- class 0.5 for train weighing

- class 1 for wagon weighing

- verification scale interval of 0.1 t

- maximum wagon weight of 90 t

The test train consists of 29 loaded wagons in total, 20 of which are reference wagons with a combined reference wagon weight of 441 t. The 20 reference wagon weights vary between 20.7 t and 23.1 t.

(a)	Wagon Weighing

The following calculations are used to determine the allowable MPE for a single reference wagon weighing 23.1 t. It is the greatest of the values calculated below, rounded to the nearest scale interval.

- The value calculated according to Table 2.

Reference wagon wt	=	23.1 t

MPE from Table 2	=	±0.5%

Mass value of MPE	=	±0.5%  23.1 t

=	±0.116 t

- The value calculated according to Table 2 for the weight of a single wagon equal to 35% of the maximum wagon weight (as marked on the data plate).

Maximum wagon wt	=	90 t

35% of max wagon wt	=	31.5 t

MPE from Table 2	=	±0.5%

Mass value of MPE	=	±0.5%  31.5 t

=	±0.157 t

- ± One scale interval.

One scale interval	=	0.1 t

Mass value of MPE	=	±0.1 t

Therefore the MPE calculated in point 2, rounded to the nearest scale interval, is the allowable MPE for the 23.1 t reference wagon, i.e. ±0.2 t. Table 9 also outlines these calculations.

Note:	You have to calculate the allowable MPE for each reference wagon.

At verification/certification, it is acceptable for up to 10% of the weighing results taken from one or more passes of the test train to exceed the appropriate MPE in Table 2 but they shall not exceed two times that value. Therefore:

- 90% of the reference wagons (18) can have a maximum error of ±0.2 t; and

- 10% of the reference wagons (2) can have a maximum error of ±0.4 t.

(b)	Train Weighing

The allowable MPE for the total weight of the reference wagons is the greatest of the values calculated below, rounded to the nearest scale interval.

- The value calculated according to Table 2.

Combined ref wgn wt	=	441 t

MPE from Table 2	=	±0.25%

Mass value of MPE 	=	±0.25%  441 t

=	±1.103 t

- The value calculated according to Table 2 of the weight of a single wagon equal to 35% of the maximum wagon weight  (as marked on the data plate) multiplied by the number of reference wagons in the test train (not exceeding 10 wagons).

Maximum wagon wt	=	90 t

35% of max wagon wt	=	31.5 t

MPE from Table 2	=	±0.25%

Mass value of MPE	=	±0.25%  31.5 t

=	±0.079 t

Multiplied by not more than 10 wagons
=	±0.079 t 10	=	±0.79 t

- ± One scale interval for each wagon in the test train (but not exceeding 10 scale intervals).

One scale interval	=	0.1 t

Mass value of MPE	=	±0.1 t  10

=	±1 t

Therefore the MPE calculated in point 1, rounded to the nearest scale interval, is the allowable MPE for the total weight of the reference wagons, 
i.e. ±1.1 t. Table 10 also outlines these calculations.

Table 9. Allowable MPE for wagon weighing (Example 2)
(greatest value below, rounded to nearest scale interval) = ±0.2 t

| MPE from Table 2  reference wagon weight = ±0.5%  23.1 t = | ±0.116 t |
| --- | --- |
| MPE from Table 2  35% of maximum wagon weight marked on the data plate = 0.5%  90 t  35/100 = | ±0.157 t |
|  one scale interval = | ±0.1 t |

Table 10. Allowable MPE for train weighing (Example 2)
(greatest value below, rounded to nearest scale interval) = ±1.1 t

| MPE from Table 2  sum of all the reference wagon weights = ±0.25%  441 t = | ±1.103 t |
| --- | --- |
| MPE from Table 2  35% of maximum wagon weight marked on the data plate  number of reference wagons (but not exceeding ten) = ±0.25%  90 t  35/100  10 = | ±0.79 t |
|  one scale interval, but not exceeding ten scale intervals = ±0.1 t  10 = | ±1 t |