Part 2: Test Report Format

NMI M 6-2
Electricity Meters

© Commonwealth of Australia 2010

First edition	—	September 2010

National Measurement Institute

Bradfield Road, Lindfield, NSW 2070

PO Box 264, Lindfield, NSW 2070

T	(61 2) 8467 3600

F	(61 2) 8467 3610

W	www.measurement.gov.au

CONTENTS

1.	Test Information	1

2.	Meter Information and Specifications	1

3.	Requirements	2

4.	Maximum Permissible Errors	4

4.1	Direct-connected Meters with Balanced Loads	4

4.2	Transformer-operated Meters with Balanced Loads	4

4.3	Polyphase Direct-connected Meters	5

4.4	Polyphase Transformer-operated Meters	5

5.	Influence Factors and Disturbances	6

5.1	Voltage Variation	6

5.2	Frequency Variation	8

5.3	Harmonic Components in the Current and Voltage Circuits	9

5.4	Reversed Phase Sequence	9

5.5	Voltage Unbalance	10

5.6	Auxiliary Voltage ±15%	10

5.7	DC Component in the AC Circuit	10

5.8	Continuous Magnetic Induction of External Origin	10

5.9	Magnetic Induction of External Origin 0.5 mT	11

5.10	Electromagnetic RF Fields	11

5.11	Conducted RF Fields	11

5.12	Fast Transient Bursts	12

5.13	Variations due to Short-time Overcurrents	12

5.14	Operation of Accessories	13

5.15	Sub-harmonics in the AC Circuit	13

5.16	Odd Harmonics in the AC Circuit	13

5.17	Tilt at 3° in any Direction from the Vertical	13

5.18	Current Coil Self-heating	14

5.19	Alternative Usage and Phase Reversal (Balanced Two-element Driven)	14

5.20	Alternative Usage and Phase Reversal (Single-element Driven)	14

5.21	Register Friction	15

5.22	Register Changeover	15

5.23	Shock	15

6.	Ambient Temperature Variation	16

6.1	Direct-connected Meters	16

6.2	Transformer-operated Meters	17

7.	Internal Clocks	17

7.1	Mains Supply	17

7.2	Operational Reserve	18

7.3	High Temperature	18

7.4	Low Temperature	18

8.	Performance Tests	18

8.1	Optical Port Requirements	18

8.2	Dry Heat Test	18

8.3	Cold Test	19

8.4	Damp Heat Cyclic Test	19

8.5	Solar Radiation Test	20

8.6	Dust Test	20

8.7	Vibration (Sinusoidal) Test	20

8.8	Mechanical Shock Test	21

8.9	Radiated Electromagnetic Radiofrequency Fields Test without Current	21

8.10	Electrostatic Discharge Test	21

8.11	Voltage Dips and Short-term Interruptions Test	22

8.12	Impulse Voltage Test	22

8.13	AC Voltage Test	22

PREFACE

This document provides a test report format for the pattern approval and verification of active-energy static electricity meters of classes 0.2, 0.5, 1 and 1.5 covered by NMI M 6-1 Electricity Meters. 
Part 1: Metrological and Technical Requirements. It may clarify NMI M 6-1, but it does not add to or alter any requirements.

The test reports should make testing more efficient and consistent, and if testing has been conducted against other standards, the test reports may be used as a checklist to determine what further testing and information is required.

Please note that not all tests apply to all meter.

# Test Information

| Test Report |  |
| --- | --- |
| Report reference number |  |
| Date of issue |  |
| Date of testing |  |
| Laboratory details |  |
| Name |  |
| Address |  |
| Contact details |  |
| Test specification |  |
| Standard | NMI M 6-1 |
| Client details |  |
| Applicant |  |
| Address |  |

# Meter Information and Specifications

| Manufacturer |  |
| --- | --- |
| Model |  |
| Accuracy class |  |
| Temperature ranges |  |
| Specified operating range |  |
| Limit range of operation |  |
| Storage and transportation |  |
| Number of phases |  |
| Number of elements |  |
| Number of wires |  |
| Connection type |  |
| Design type (e.g. solid state, induction) |  |
| Reference frequency |  |
| Reference voltage |  |
| Basic current (direct-connected) |  |
| Rated current (transformer-operated) |  |
| Rated maximum current |  |
| Clock type |  |
| Enclosure type |  |
| Protective class |  |
| Software/firmware version |  |

# Requirements

| Clause number and requirement (from NMI M 6-1) |  |  | Value / Remark |  | Result |
| --- | --- | --- | --- | --- | --- |
| 3 | Units of measurement |  |  |  |  |
|  | Valid units of measurement used |  |  |  |  |
| 4.1 | Minimum measured quantity |  |  |  |  |
|  | Has the form 1 × 10n authorised units of energy, where n is an integer |  |  |  |  |
| 4.2 | Maximum permissible variation between indicators |  |  |  |  |
|  | No indicated difference between indications of same quantity on different indicators |  |  |  |  |
| 4.3 | Calculated quantities |  |  |  |  |
|  | Indicated quantity equals value obtained using indicated values with applicable rounding |  |  |  |  |
|  | If rounding applied it is 0.5 minimum measured quantity |  |  |  |  |
| 4.6 | Meter constant |  |  |  |  |
|  | No error in relationship between test output and indication on display |  |  |  |  |
| 4.7 | Class indices (accuracy class) |  |  |  |  |
|  | Meter classified as one of 0.2, 0.5, 1 or 1.5 |  |  |  |  |
| 4.8 | Maximum permissible error |  |  |  |  |
|  | Percentage errors do not exceed the relevant values specified in tables 1, 2 and 3 due to variations in current |  | (see results below) |  |  |
| 5.2 | Temperature range (ranges shall comply with the minimum acceptable ranges in Table 5) |  |  |  |  |
|  | Specified operating range |  |  |  |  |
|  | Limit range of operation |  |  |  |  |
|  | Storage and transportation |  |  |  |  |
| 5.7.2 |  | Initial start up of the meter |  |  |  |
|  |  | Time to start – shall be functional within 5 s |  |  |  |
| 5.7.3 |  | Running with no load |  |  |  |
|  |  | Test voltage |  |  |  |
|  |  | Test current |  |  |  |
|  |  | Test period |  |  |  |
|  |  | Test output pulses – shall be no more than one |  |  |  |
|  |  | Rotor revolutions – may start but shall not complete a revolution |  |  |  |
| 5.7.4 |  | Starting |  |  |  |
|  |  | Test current |  |  |  |
|  |  | Meter starts and continues to register |  |  |  |
|  |  | Rotor revolutions – shall start and complete at least one revolution |  |  |  |
| 7.2 |  | Acting upon significant faults (static meters only) |  |  |  |
|  |  | Has capability to detect, log and communicate significant faults |  |  |  |
|  |  | Logged data kept in permanent record with date and time stamp |  |  |  |
| 7.3 |  | Display |  |  |  |
|  |  | Meter has a display which is legible whilst operating |  |  |  |
|  |  | Visible to consumer in normal installation position |  |  |  |
|  |  | There is a procedure to show all relevant elements of indicator display, with sufficient time to check them |  |  |  |
|  |  | Able to display quantity of energy corresponding to Imax for at least 4 000 h without returning to same index |  |  |  |
|  |  | Calculated value (energy at Imax for 4 000 h) |  |  |  |
|  |  | Number of display digits |  |  |  |
| 7.4 |  | Auxiliary devices interface |  |  |  |
|  |  | Interface shall be sealed if parameters can be altered by instructions or data introduced through interface |  |  |  |
| 8.1 |  | Information to be displayed on meter exterior |  |  |  |
|  |  | Manufacturer’s name or mark |  |  |  |
|  |  | Model designation |  |  |  |
|  |  | Serial number |  |  |  |
|  |  | NMI certificate of approval number (space for) |  |  |  |
|  |  | Number of phases, number of wires |  |  |  |
|  |  | Reference frequency |  |  |  |
|  |  | Specified operating temperature limits (if more restrictive than –10°C to +60°C) |  |  |  |
|  |  | Meter constant |  |  |  |
|  |  | Rated voltage |  |  |  |
|  |  | Rated currents |  |  |  |
|  |  | Class index |  |  |  |
| 8.2 |  | Notices |  |  |  |
|  |  | Any special notices or limitations of use shall be clearly marked or provided in manual |  |  |  |
| 9.1 |  | Verification mark |  |  |  |
|  |  | Easily affixed without affecting metrological properties of the meter |  |  |  |
|  |  | Visible without moving or dismantling the meter when in use |  |  |  |
|  |  | Part where mark is located shall not be removable with damage to mark |  |  |  |
|  |  | Sufficient space (≥ 200 mm2) |  |  |  |
| 9.2 |  | Sealing |  |  |  |
|  |  | Access to protected parameters protected |  |  |  |
|  |  | Access to protected parameters recorded |  |  |  |
|  |  | Records readily accessible |  |  |  |
|  |  | Record easily identifiable (not confused) |  |  |  |
|  |  | Reference record marked on meter |  |  |  |
|  |  | Record shall not repeat in a sequence of less than 99 alterations; record shall persist reliably for at least 2 years and persist through influence and disturbance tests |  |  |  |

# Maximum Permissible Errors

Refer to NMI M 6-1, clause 4.8 (Tables 1 to 3).

Mandatory test currents are shown. Extra test points should be included where appropriate for the particular meter.

## Direct-connected Meters with Balanced Loads

| Current (A) | Power factor | Percentage error | Limit (±%) for class |  |
| --- | --- | --- | --- | --- |
|  |  |  | 1 | 1.5 |
| 0.05 Ib | 1 |  | 1.5 | 1.5 |
| 0.1 Ib |  |  | 1.0 | 1.5 |
| 0.2 Ib |  |  |  |  |
| Ib |  |  |  |  |
| Imax |  |  |  |  |
| 0.1 Ib | 0.5 inductive |  | 1.5 | 1.5 |
| 0.2 Ib |  |  | 1.0 | 1.5 |
| Ib |  |  |  |  |
| Imax |  |  |  |  |
| 0.1 Ib | 0.8 capacitive |  | 1.5 |  |
| 0.2 Ib |  |  | 1.0 |  |
| Ib |  |  |  |  |
| Imax |  |  |  |  |

## Transformer-operated Meters with Balanced Loads

| Current (A) | Power factor | Percentage error | Limit (±%) for class |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 0.2 | 0.5 | 1 |
| 0.01 In | 1 |  | 0.4 | 1.0 |  |
| 0.02 In |  |  |  |  | 1.5 |
| 0.05 In |  |  | 0.2 | 0.5 | 1.0 |
| 0.1 In |  |  |  |  |  |
| In |  |  |  |  |  |
| Imax |  |  |  |  |  |
| 0.02 In | 0.5 inductive |  | 0.5 | 1.0 |  |
| 0.05 In |  |  |  |  | 1.5 |
| 0.1 In |  |  | 0.3 | 0.6 | 1.0 |
| In |  |  |  |  |  |
| Imax |  |  |  |  |  |
| 0.02 In | 0.8 capacitive |  | 0.5 | 1.0 |  |
| 0.05 In |  |  |  |  | 1.5 |
| 0.1 In |  |  | 0.3 | 0.6 | 1.0 |
| In |  |  |  |  |  |
| Imax |  |  |  |  |  |

## Polyphase Direct-connected Meters

| Phase | Current (A) | Power factor | Percentage error | Limit (±%) for class |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | 1 | 1.5 |
| Phase 1 | 0.1 Ib | 1 |  | 2.0 | 2.5 |
|  | 0.2 Ib |  |  |  |  |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.2 Ib | 0.5 inductive |  | 2.0 | 2.5 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| Phase 2 | 0.1 Ib | 1 |  | 2.0 | 2.5 |
|  | 0.2 Ib |  |  |  |  |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.2 Ib | 0.5 inductive |  | 2.0 | 2.5 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| Phase 3 | 0.1 Ib | 1 |  | 2.0 | 2.5 |
|  | 0.2 Ib |  |  |  |  |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.2 Ib | 0.5 inductive |  | 2.0 | 2.5 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |

## Polyphase Transformer-operated Meters

| Phase | Current (A) | Power factor | Percentage error | Limit (±%) for class |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 | 1 |
| Phase 1 | 0.05 In | 1 |  | 0.3 | 0.6 | 2.0 |
|  | 0.1 In |  |  |  |  |  |
|  | In |  |  |  |  |  |
|  | Imax |  |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.4 | 1.0 | 2.0 |
|  | In |  |  |  |  |  |
|  | Imax |  |  |  |  |  |
| Phase 2 | 0.05 In | 1 |  | 0.3 | 0.6 | 2.0 |
|  | 0.1 In |  |  |  |  |  |
|  | In |  |  |  |  |  |
|  | Imax |  |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.4 | 1.0 | 2.0 |
|  | In |  |  |  |  |  |
|  | Imax |  |  |  |  |  |
| Phase 3 | 0.05 In | 1 |  | 0.3 | 0.6 | 2.0 |
|  | 0.1 In |  |  |  |  |  |
|  | In |  |  |  |  |  |
|  | Imax |  |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.4 | 1.0 | 2.0 |
|  | In |  |  |  |  |  |
|  | Imax |  |  |  |  |  |

# Influence Factors and Disturbances

## Voltage Variation

Refer to NMI M 6-1, Table 4 and A.2.12.

For three phase mains power, voltage variations shall apply for each phase successively.

    - Direct-connected Meters, Classes 1 and 1.5

| Voltage variation (% from Unom) | Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | 1 | 1.5 |
| +10 | 0.05 Ib | 1 |  | 0.7 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  | 1.0 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| –10 | 0.05 Ib | 1 |  | 0.7 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  | 1.0 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| +15 | 0.05 Ib | 1 |  | 2.1 | 3.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  | 3.0 | 3.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| –20 | 0.05 Ib | 1 |  | 2.1 | 3.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  | 3.0 | 3.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| –50 | 0.05 Ib | 1 |  | –100 to +10 |  |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  |  |  |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |

    - Transformer-operated Meters, Class 1

| Voltage variation (% from Unom) | Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- |
|  |  |  |  | 1 |
| +10 | 0.02 In | 1 |  | 0.7 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  | 1.0 |
|  | In |  |  |  |
|  | Imax |  |  |  |
| –10 | 0.02 In | 1 |  | 0.7 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  | 1.0 |
|  | In |  |  |  |
|  | Imax |  |  |  |
| +15 | 0.02 In | 1 |  | 2.1 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  | 3.0 |
|  | In |  |  |  |
|  | Imax |  |  |  |
| –20 | 0.02 In | 1 |  | 2.1 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  | 3.0 |
|  | In |  |  |  |
|  | Imax |  |  |  |
| –50 | 0.02 In | 1 |  | –100 to +10 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  |  |
|  | In |  |  |  |
|  | Imax |  |  |  |

    - Transformer-operated Meters, Classes 0.2 and 0.5

| Voltage variation (% from Unom) | Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 |
| +10 | 0.05 In | 1 |  | 0.1 | 0.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.2 | 0.4 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
| –10 | 0.05 In | 1 |  | 0.1 | 0.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.2 | 0.4 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
| +15 | 0.05 In | 1 |  | 0.3 | 0.6 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.6 | 1.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
| –20 | 0.05 In | 1 |  | 0.3 | 0.6 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.6 | 1.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
| –50 | 0.05 In | 1 |  | –100 to +10 |  |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  |  |  |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |

## Frequency Variation

Refer to NMI M 6-1, Table 4 and A.2.13.

    - Direct-connected meters, Classes 1 and 1.5

| Frequency variation (% from fnom) | Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | 1 | 1.5 |
| +2 | 0.05 Ib | 1 |  | 0.5 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  | 0.7 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
| –2 | 0.05 Ib | 1 |  | 0.5 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 Ib | 0.5 inductive |  | 0.7 | 1.0 |
|  | Ib |  |  |  |  |
|  | Imax |  |  |  |  |

    - Transformer-operated Meters, Class 1

| Frequency variation (% from fnom) | Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- |
|  |  |  |  | 1 |
| +2 | 0.02 In | 1 |  | 0.5 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  | 0.7 |
|  | In |  |  |  |
|  | Imax |  |  |  |
| –2 | 0.02 In | 1 |  | 0.5 |
|  | In |  |  |  |
|  | Imax |  |  |  |
|  | 0.05 In | 0.5 inductive |  | 0.7 |
|  | In |  |  |  |
|  | Imax |  |  |  |

    - Transformer-operated Meters, Classes 0.2 and 0.5

| Frequency variation (% from fnom) | Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 |
| +2 | 0.05 In | 1 |  | 0.1 | 0.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.1 | 0.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
| –2 | 0.05 In | 1 |  | 0.1 | 0.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |
|  | 0.1 In | 0.5 inductive |  | 0.1 | 0.2 |
|  | In |  |  |  |  |
|  | Imax |  |  |  |  |

## Harmonic Components in the Current and Voltage Circuits

Refer to NMI M 6-1, Table 4.

The variation in percentage error shall be measured under the most unfavourable phase displacement of the fifth harmonic in the current compared with the fundamental error.

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | fnom | fnom + harmonics |  | 0.2 | 0.5 | 1 | 1.5 |
| 0.5 Imax | 1 |  |  |  | 0.4 | 0.5 | 0.8 | 1.0 |

## Reversed Phase Sequence

Refer to NMI M 6-1, Table 4.

Polyphase (three-phase four wire) meters shall measure and register within the limits of variation in percentage error if any one or two phases of the three phase network are interrupted.

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | ABC | CBA |  | 0.2 | 0.5 | 1 | 1.5 |
| 0.1 Ib (0.1 In) | 1 |  |  |  | 0.05 | 0.1 | 1.5 | 1.5 |

## Voltage Unbalance

Refer to NMI M 6-1, Table 4.

Polyphase (three-phase four wire) meters shall measure and register within the limits of variation in percentage error if any one or two phases of the three phase network are interrupted.

| Current (A) | Power factor | Phases interrupted | Variation in error (%) | Limit of variation (%) by class |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 | 1 |
| Ib (In) | 1 | 1 phase – A |  | 0.5 | 1.0 | 2.0 |
|  |  | 1 phase – B |  |  |  |  |
|  |  | 1 phase – C |  |  |  |  |
|  |  | 2 phases – AB |  |  |  |  |
|  |  | 2 phases – AC |  |  |  |  |
|  |  | 2 phases – BC |  |  |  |  |

## Auxiliary Voltage ±15%

Refer to NMI M 6-1, Table 4.

Applicable only if the auxiliary supply is not internally connected to the voltage measuring circuit.

| Voltage (% from reference) | Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Reference | Measured |  | 0.2 | 0.5 |
| +15 | 0.01 In | 1 |  |  |  | 0.05 | 0.1 |
| –15 |  |  |  |  |  |  |  |

## DC Component in the AC Circuit

Refer to NMI M 6-1, Table 4.

This test does not apply to transformer-operated meters.

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | fnom | + DC component |  | 1 | 1.5 |
| Imax/√2 | 1 |  |  |  | 3.0 | 6.0 |

## Continuous Magnetic Induction of External Origin

Refer to NMI M 6-1, Table 4.

| Current (A) | Power factor | Position of magnet | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 | Front |  | 2.0 | 2.0 | 2.0 | 3.0 |
|  |  | Left-hand side |  |  |  |  |  |
|  |  | Right-hand side |  |  |  |  |  |
|  |  | Top |  |  |  |  |  |
|  |  | Bottom |  |  |  |  |  |

## Magnetic Induction of External Origin 0.5 mT

Refer to NMI M 6-1, Table 4.

A magnetic induction of external origin of 0.5 mT produced by a current of the same frequency as that of the voltage applied to the meter and under the most unfavourable conditions of phase and direction shall not cause a variation in the percentage error of the meter exceeding the values shown. The magnetic induction shall be obtained by placing the meter in the centre of a circular coil, 1 m in mean diameter, of square section and of small radial thickness relative to the diameter, and having 400 At.

| Current (A) | Power factor | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  | 0.5 | 1.0 | 2.0 | – |
| Ib (0.5 In) | 1 |  | – | – | – | 2.0 |

## Electromagnetic RF Fields

Refer to NMI M 6-1, Table 4 and A.2.9 (test with current test).

Meters constructed with passive elements only, including electromechanical meters, are exempt from this test.

Frequency range:	0 to 2 400 MHz

Modulation:	80% AM, 1kHz sine wave

Field strength:	10 V/m

| Current (A) | Power factor | Polarisation | Facing meter | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 | Vertical | Front |  | 1.0 | 2.0 | 2.0 | 3.0 |
|  |  |  | Right |  |  |  |  |  |
|  |  |  | Left |  |  |  |  |  |
|  |  |  | Rear |  |  |  |  |  |
|  |  | Horizontal | Front |  |  |  |  |  |
|  |  |  | Right |  |  |  |  |  |
|  |  |  | Left |  |  |  |  |  |
|  |  |  | Rear |  |  |  |  |  |
| Requirement |  |  |  | Remark |  |  |  | Result |
| During the test, the behaviour of the meter shall not be perturbed |  |  |  |  |  |  |  |  |

## Conducted RF Fields

Refer to NMI M 6-1, Table 4 and A.2.10.

Meters constructed with passive elements only, including electromechanical meters, are exempt from this test.

RF amplitude (50 Ω):	10 V (e.m.f.)

Modulation:	80% AM, 1 kHz sine wave

Frequency range:	0.15 to 80 MHz

| Current (A) | Power factor | Power port or I/O port | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  |  | 1.0 | 2.0 | 2.0 | 3.0 |
| Requirement |  |  | Remark |  |  |  | Result |
| During the test, the behaviour of the meter shall not be perturbed |  |  |  |  |  |  |  |

## Fast Transient Bursts

Refer to NMI M 6-1, Table 4 and A.2.15.

Meters constructed with passive elements only, including electromechanical meters, are exempt from this test. During the test, a temporary degradation or loss of function or performance is acceptable.

| Current (A) | Power factor | Circuit | Voltage peak (kV) | Polarity (60 s at each) | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 | Voltage | 4 | Positive |  | 1.0 | 2.0 | 4.0 | 6.0 |
|  |  |  |  | Negative |  |  |  |  |  |
|  |  | Current |  | Positive |  |  |  |  |  |
|  |  |  |  | Negative |  |  |  |  |  |
|  |  | Auxiliary circuit | 2 | Positive |  |  |  |  |  |
|  |  |  |  | Negative |  |  |  |  |  |
|  |  | Auxiliary circuit |  | Positive |  |  |  |  |  |
|  |  |  |  | Negative |  |  |  |  |  |

## Variations due to Short-time Overcurrents

Refer to NMI M 6-1, Table 4 and A.2.16.

The test shall be performed for polyphase meters phase-by-phase.

| Current (A) | Power factor | Test | Over-current value (A) | Duration (ms) | Phase | Variation in error (%) | Limit of variation (%) by class |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | 0.2 | 0.5 |  | 1 | 1.5 |
| Ib | 1 | A | 30 Imax | 10 | 1 |  | – | – |  | 1.5 | 1.5 |
|  |  |  |  |  | 2 |  |  |  |  |  |  |
|  |  |  |  |  | 3 |  |  |  |  |  |  |
| In | 1 | B | 20 Imax | 500 | 1 |  | 0.05 | 0.05 |  | 0.5 | – |
|  |  |  |  |  | 2 |  |  |  |  |  |  |
|  |  |  |  |  | 3 |  |  |  |  |  |  |
| Ib | 1 | C | 7000 | 60 | 1 |  | Requirement Meter shall not cause damage to surrounding equipment |  |  |  |  |
|  |  |  |  |  | 2 |  |  |  |  |  |  |
|  |  |  |  |  | 3 |  |  |  |  |  |  |
| In | 1 | D | 250 | 60 | 1 |  |  |  |  |  |  |
|  |  |  |  |  | 2 |  |  |  |  |  |  |
|  |  |  |  |  | 3 |  |  |  |  |  |  |
| In | 1 | E | 50 | 60 | 1 |  |  |  |  |  |  |
|  |  |  |  |  | 2 |  |  |  |  |  |  |
|  |  |  |  |  | 3 |  |  |  |  |  |  |
| Requirement |  |  |  |  | Remark |  |  |  | Result |  |  |
| For tests C, D and E the meter shall not cause damage to surrounding equipment |  |  |  |  |  |  |  |  |  |  |  |

## Operation of Accessories

Refer to NMI M 6-1, Table 4.

Such an accessory, when enclosed in the meter case, is energised intermittently, for example the electromagnet of a multi-rate register. It is preferable that the connection to the auxiliary device(s) is marked to indicate the correct method of connection. If these connections are made by means of plugs and sockets, they should be irreversible. However, in the absence of those markings or irreversible connections, the variations of errors shall not exceed those indicated in this table if the meter is tested with the connections giving the most unfavourable condition.

| Current (A) | Power factor | Accessory | Connection | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| 0.05 Ib (0.05 In) | 1 |  |  |  | – | – | 0.5 | – |
| 0.01 In | 1 |  |  |  | 0.05 | 0.1 | – | – |

## Sub-harmonics in the AC Circuit

Refer to NMI M 6-1, Table 4 and A.2.17.

Test waveform:	sinusoid, 2 cycles on, 2 cycles off

Current amplitude:	2  reference current

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | fnom | Test waveform |  | 0.2 | 0.5 | 1 | 1.5 |
| 0.5 Ib (0.5 In) | 1 |  |  |  | 0.5 | 0.75 | 1.5 | 3.0 |

## Odd Harmonics in the AC Circuit

Refer to NMI M 6-1, Table 4 and A.2.18.

Test waveform:	sinusoid, set to zero for first and third quarters of each period

Current amplitude:	2  reference current

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | fnom | Test waveform |  | 0.2 | 0.5 | 1 | 1.5 |
| 0.5 Ib (0.5 In) | 1 |  |  |  | 0.4 | 0.5 | 0.8 | 1.0 |

## Tilt at 3° in any Direction from the Vertical

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position.

| Current (A) | Power factor | Direction of 3° tilt | Variation in error (%) | Limit of variation (%) by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| 0.1 Ib | 1 | Forward |  | 1.0 | 2.0 | 2.0 | 3.0 |
|  |  | Backward |  |  |  |  |  |
|  |  | Left |  |  |  |  |  |
|  |  | Right |  |  |  |  |  |

## Current Coil Self-heating

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position.

Initial error:	determine for each load while current coil is still unheated

Precondtioning:	1 h with voltage circuit at Unom, current circuit with zero current

Test:	continue until error becomes constant

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- | --- |
|  |  | Coil unheated (initial error) | Coil heated |  | 1.5 |
| Imax | 1 |  |  |  | 1.0 |
| Imax | 0.5 inductive |  |  |  | 1.0 |

## Alternative Usage and Phase Reversal (Balanced Two-element Driven)

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position.

| Current (A) | Power factor | Phase sequence | Percentage error | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 1.5 |
| 0.05 Ib | 1 | A1 leading A2 by 180° |  |  |  |
|  |  | A1 leading A2 by 120° |  |  | 1.0 |
|  |  | A2 leading A1 by 120° |  |  |  |
| Ib | 1 | A1 leading A2 by 180° |  |  |  |
|  |  | A1 leading A2 by 120° |  |  | 0.5 |
|  |  | A2 leading A1 by 120° |  |  |  |

## Alternative Usage and Phase Reversal (Single-element Driven)

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position.

| Current (A) | Power factor | Phase sequence | Percentage error | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 1.5 |
| 0.1 Ib | 1 | A1 leading A2 by 180° |  |  |  |
|  |  | A1 leading A2 by 120° |  |  | 1.0 |
|  |  | A2 leading A1 by 120° |  |  |  |
| 2 Ib | 1 | A1 leading A2 by 180° |  |  |  |
|  |  | A1 leading A2 by 120° |  |  | 0.5 |
|  |  | A2 leading A1 by 120° |  |  |  |

## Register Friction

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position. For a multi-rate meter, the changeover device shall be in each operating condition in turn.

| Current (A) | Power factor | Changeover device operating condition | Rotor Speed |  | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Heaviest load | Register disengaged |  | 1.5 |
| 0.05 Ib | 1 |  |  |  |  | 0.5 |

## Register Changeover

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position.

| Current (A) | Power factor | Changeover device operating condition | Rotor Speed | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 1.5 |
| 0.05 Ib | 1 |  |  |  |  |
|  |  |  |  |  | 0.4* |

*	An additional variation of 0.5% may be permitted for certain multiple-element meters 
(refer to AS 1284.1, clause 4.3.15).

## Shock

Refer to NMI M 6-1, Table 4.

This test is only required for induction meters and any other meters which may be influenced by their working position.

Initial error:	determine for each load prior to subjecting to shock

| Current (A) | Power factor | Initial error (before shock) | After shock test | Variation in error (%) | Limit of variation (%) by class |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 1.5 |
| 0.05 Ib | 1 |  |  |  | 0.5 |
| Ib | 1 |  |  |  | 0.3 |
| Ib | 0.5 inductive |  |  |  | 0.3 |

# Ambient Temperature Variation

Refer to NMI M 6-1, Table 6 and A.2.3.

The meter error shall be determined at a minimum of four temperature values across the whole operating range.

## Direct-connected Meters

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Mean temperature coefficient (%/K) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | TLow | THigh |  | Calculated | Limit by class |  |
|  |  |  |  |  |  | 1 | 1.5 |
| Temperature interval, TLow to THigh (e.g. –10°C to 15°C) |  |  |  |  |  |  |  |
| 0.1 Ib | 1 |  |  |  |  | 0.05 | 0.05 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| 0.2 Ib | 0.5 inductive |  |  |  |  | 0.07 | 0.07 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| Temperature interval, TLow to THigh (e.g. 5°C to 25°C) |  |  |  |  |  |  |  |
| 0.1 Ib | 1 |  |  |  |  | 0.05 | 0.05 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| 0.2 Ib | 0.5 inductive |  |  |  |  | 0.07 | 0.07 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| Temperature interval, TLow to THigh (e.g. 25°C to 45°C) |  |  |  |  |  |  |  |
| 0.1 Ib | 1 |  |  |  |  | 0.05 | 0.05 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| 0.2 Ib | 0.5 inductive |  |  |  |  | 0.07 | 0.07 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |

## Transformer-operated Meters

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Mean temperature coefficient (%/K) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | TLow | THigh |  | Calculated | Limit by class |  |
|  |  |  |  |  |  | 1 | 1.5 |
| Temperature interval, TLow to THigh (e.g. –10°C to 15°C) |  |  |  |  |  |  |  |
| 0.1 Ib | 1 |  |  |  |  | 0.05 | 0.05 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| 0.2 Ib | 0.5 inductive |  |  |  |  | 0.07 | 0.07 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| Temperature interval, TLow to THigh (e.g. 5°C to 25°C) |  |  |  |  |  |  |  |
| 0.1 Ib | 1 |  |  |  |  | 0.05 | 0.05 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| 0.2 Ib | 0.5 inductive |  |  |  |  | 0.07 | 0.07 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| Temperature interval, TLow to THigh (e.g. 25°C to 45°C) |  |  |  |  |  |  |  |
| 0.1 Ib | 1 |  |  |  |  | 0.05 | 0.05 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |
| 0.2 Ib | 0.5 inductive |  |  |  |  | 0.07 | 0.07 |
| Ib |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |

# Internal Clocks

Refer to NMI M 6-1, clause 6.

This test applies to any solid state internal clock used for electricity meters and load control devices.

Internal clock type (synchronous or crystal-controlled)

Operational reserve (spring or battery/super-capacitor/primary cell)

## Mains Supply

Testing period:	30 days

Test temperature:	23°C

|  | Date | Time |  | Difference (s) | Limit variation (s/day) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Ref | Test |  | Result | Synchronous | Crystal |
| Start |  |  |  |  |  | 0.167 | 0.5 |
| End |  |  |  |  |  |  |  |

## Operational Reserve

Testing period:	36 h

Test temperature:	23°C

|  | Date | Time |  | Difference (s) | Limit variation (s/day) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Ref | Test |  | Result | Synchronous |  | Crystal |
|  |  |  |  |  |  | Spring | Battery |  |
| Start |  |  |  |  |  | 120 | 1 | 1 |
| End |  |  |  |  |  |  |  |  |

## High Temperature

Testing period:	24 h

Test temperature:	45°C

|  | Date | Time |  | Difference (s) | Limit variation (s/day) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Ref | Test |  | Result | Crystal |
| Start |  |  |  |  |  | 0.15 |
| End |  |  |  |  |  |  |

## Low Temperature

Testing period:	24 h

Test temperature:	–10°C

|  | Date | Time |  | Difference (s) | Limit variation (s/day) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Ref | Test |  | Result | Crystal |
| Start |  |  |  |  |  | 0.15 |
| End |  |  |  |  |  |  |

# Performance Tests

## Optical Port Requirements

Refer to NMI M 6-1, A.1.3.

| Requirement | Remark | Result |
| --- | --- | --- |
| Environmental lighting condition |  |  |
| Transmission speed |  |  |

## Dry Heat Test

Refer to NMI M 6-1, A.2.1.

Duration:	72 h

Meter/EUT:	In operating condition except whilst temperature is lowered or raised.

High temperature:		(maximum specified operating temperature)

| Current (A) | Power factor | Percentage error |  |  | MPE by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | At reference before heat | At high temperature after 72 h | Reference after recovery | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  |  | Remark |  |  |  | Result |  |
| No damage to meter |  |  |  |  |  |  |  |  |
| No change of information |  |  |  |  |  |  |  |  |

## Cold Test

Refer to NMI M 6-1, A.2.2.

Duration:	72 h

Meter/EUT:	in operating condition except whilst temperature is lowered or raised.

Low temperature:		(minimum specified operating temperature)

| Current (A) | Power factor | Percentage error |  |  | MPE by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | At reference before heat | At low temperature after 72 h | Reference after recovery | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  |  | Remark |  |  |  | Result |  |
| No damage to meter |  |  |  |  |  |  |  |  |
| No change of information |  |  |  |  |  |  |  |  |

## Damp Heat Cyclic Test

Refer to NMI M 6-1, A.2.4.

Duration (cycles):	6  24 h cycles

Meter/EUT:	non-operating condition

Low temperature:	25°C

High temperature:		(maximum specified operating temperature)

| Requirement | Remark | Result |
| --- | --- | --- |
| No trace of corrosion likely to affect the functional properties of the EUT shall be present |  |  |

24 h after the end of this test, submit the EUT to the following tests:

- AC voltage test (NMI M 6-1, A.2.20) — refer to clause 8.13.

- impulse voltage test (NMI M 6-1, A.2.19, except 0.8 of impulse voltage, i.e. 9.6 kV +0%, –15%)

| Impulse voltage test |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit |  |
|  |  | Before | After |  |  |  |
| Ib (In) | 1 |  |  |  | (uncertainty of measurement) |  |
| Requirement |  |  |  | Remark |  | Result |
| During the test, no flashover, disruptive discharge or puncture shall occur |  |  |  |  |  |  |
| After the test, no mechanical damage to the EUT |  |  |  |  |  |  |

## Solar Radiation Test

Refer to NMI M 6-1, A.2.5.

UV lamp output:	21 750 lm to 27 000 lm

Duration:	48 h and distance of 250 mm

Meter/EUT:	non-operating condition

| Current (A) | Power factor | Percentage error | MPE by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  | Remark |  |  | Result |  |
| For transparent parts – no noticeable deterioration or loss in transparency |  |  |  |  |  |  |
| For non-transparent parts – no noticeable effect |  |  |  |  |  |  |
| The function of the meter shall not be impaired (see error above) |  |  |  |  |  |  |

## Dust Test

Refer to NMI M 6-1, A.2.6.

Enclosure category:	2

Duration:	8 h

Meter/EUT:	non-operating condition

| Current (A) | Power factor | Percentage error | MPE by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  | Remark |  |  | Result |  |
| No dust accumulation which could affect meter operation or safety |  |  |  |  |  |  |
| No dust deposition that could lead to tracking along creepage distances |  |  |  |  |  |  |
| The function of the meter shall not be impaired (see error above) |  |  |  |  |  |  |

## Vibration (Sinusoidal) Test

Refer to NMI M 6-1, A.2.7.

Severity level:	2

Frequency range:	10 to 150 Hz

Max acceleration level:	10 m/s2

No sweep cycles per axis:	10

Meter/EUT:	non-operating condition

| Current (A) | Power factor | Percentage error | MPE by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  | Remark |  |  | Result |  |
| No damage to meter |  |  |  |  |  |  |
| No change of information |  |  |  |  |  |  |
| Meter shall operate correctly (see error above) |  |  |  |  |  |  |

## Mechanical Shock Test

Refer to NMI M 6-1, A.2.8.

Severity level:	1

Pulse shape:	half-sine

Peak acceleration:	200 m/s2

Pulse duration:	18 ms

Meter/EUT:	non-operating condition, without packing

| Current (A) | Power factor | Percentage error | MPE by class |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  | Remark |  |  | Result |  |
| No damage to meter |  |  |  |  |  |  |
| No change of information |  |  |  |  |  |  |
| Meter shall operate correctly (see error above) |  |  |  |  |  |  |

## Radiated Electromagnetic Radiofrequency Fields Test without Current

Refer to NMI M 6-1, A.2.9.

Frequency range:	80 to 2400 MHz (continuous)

Modulation:	80% AM, 1 kHz sine wave

Field strength:	30 V/m

Meter/EUT:	in operating condition, reference voltage, current terminal open-circuit

| Requirement | Remark | Result |
| --- | --- | --- |
| The behaviour of the equipment shall not be perturbed |  |  |

## Electrostatic Discharge Test

Refer to NMI M 6-1, A.2.11.

Number of discharges:	at least 10

Polarity of discharges:	the most sensitive polarity

Severity level:	4

Meter/EUT:	in operating condition, reference voltage, current terminal open-circuit

| Application (direct/indirect) | Coupling plane | Discharge mode (contact/air) | Test voltage (kV) | Change in |  |  | Limit, x (kW·h) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Register | Test output |  |  |  |
| Direct | – |  |  |  |  |  |  |  |
| Indirect | Horizontal | Contact | 15 |  |  |  |  |  |
| Indirect | Vertical | Contact | 15 |  |  |  |  |  |
| Current (A) | Power factor |  | Percentage error |  | MPE by class |  |  |  |
|  |  |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Ib (In) | 1 |  |  |  | 0.2 | 0.5 | 1 | 1.5 |
| Requirement |  |  | Remark |  |  |  | Result |  |
| Meter shall operate correctly (see error above) |  |  |  |  |  |  |  |  |

## Voltage Dips and Short-term Interruptions Test

Refer to NMI M 6-1, A.2.14.

Meter/EUT:	in operating condition, reference voltage, no current

| Voltage dip or interruption | U | Duration | Dips/interruptions |  | Change in |  | Limit, x (kW·h) |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Number | Time between | Register | Test output |  |
| Dip | 50% | 1 min | 1 | – |  |  |  |
| Interruption | 100% | 1 s | 3 | 50 ms |  |  |  |
| Interruption | 100% | 20 ms | 1 | – |  |  |  |

## Impulse Voltage Test

Refer to NMI M 6-1, A.2.19.

Impulse voltage:	12 kV +0%, –15%

Source capacitance:	0.125 μF

Source impedance:	40 Ω ± 5 Ω

Stored energy:	9.0 J ± 1.0 J

Impulse waveform at no load:	1.2/50 impulse

Meter/EUT:	non-operating condition

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Before | After |  |  |  |
| Ib (In) | 1 |  |  |  | (uncertainty of measurement) |  |
| Requirement |  |  |  | Remark |  | Result |
| During the test, no flashover, disruptive discharge or puncture shall occur |  |  |  |  |  |  |
| After the test, no mechanical damage to the EUT |  |  |  |  |  |  |

## AC Voltage Test

Refer to NMI M 6-1, A.2.20.

This test shall be performed as part of the damp heat cyclic test (refer to NMI M 6-1, A.2.4).

| Current (A) | Power factor | Percentage error |  | Variation in error (%) | Limit |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Before | After |  |  |  |
| Ib (In) | 1 |  |  |  | (uncertainty of measurement) |  |
| Requirement |  |  |  | Remark |  | Result |
| 2 kV: during the test, no flashover, disruptive discharge or puncture shall occur |  |  |  |  |  |  |
| 4 kV: during the test, no flashover, disruptive discharge or puncture shall occur |  |  |  |  |  |  |
| 40 V: during the test, no flashover, disruptive discharge or puncture shall occur |  |  |  |  |  |  |
| After the test, no mechanical damage to the EUT |  |  |  |  |  |  |