- NMI P 105

Pattern Approval Specifications for the 
Graduation of Analogue Scales

First edition — May 1976 (Document 105)

First edition, first revision — May 1979 (Document 105)

First edition, second revision — November 1980 (Document 105)

First edition, third revision — August 1988 (Document 105)

First edition, fourth revision — June 2001 (renamed NSC P 105)

First edition, fifth revision — July 2004 (renamed NMI P 105)

Bradfield Road, Lindfield, NSW 2070

PO Box 264, Lindfield, NSW 2070

Telephone:	(61 2) 8467 3600

Facsimile:	(61 2) 8467 3610

Web page:	http://www.measurement.gov.au

© Commonwealth of Australia 1976

CONTENTS

1.	Form of Scale Marks	1

2.	Arrangement of Scale Marks	1

3.	Numbering of Scale Marks	1

4.	Recommended Numbering Arrangements	1

4.1	First Alternative	1

4.2	Second Alternative	2

4.3	Third Alternative	2

# 1.	Form of Scale Marks

Scale marks shall be made up of lines of the same thickness and the thickness shall be constant and shall be between one-tenth and one-quarter the width of a scale spacing without being less then 0.2 mm.

The length of the shortest lines shall not be less than the width of a scale spacing.

# 2.	Arrangement of Scale Marks

Three recommended arrangements for graduated scales are as follows, where n is any positive or negative whole number or zero:

- style A (top), numerical value of graduation, d, is 1×10n;

- style B (middle) numerical value of graduation, d, is 2×10n;

- style C (bottom) numerical value of graduation, d, is 5×10n.

Special consideration will be given to alternative arrangements if space or other restrictions prevent the recommended arrangements being met.

# 3.	Numbering of Scale Marks

- The actual or apparent height of numbers shall be not less then L/330, without being less than 2 mm, where L is the reading distance.

- The width of a number measured parallel to the base of the scale shall be less than the distance between two consecutive numbered scale marks.

- The decimal marker on numbers less than unity shall be preceded by a zero digit.

- The decimal marker shall be a point with a diameter of 1.5 times the thickness of the lines forming adjacent digits; the point shall be on the base line of the digits or at mid-height.

- Digits of the number shall be horizontally aligned except on fan-shaped dials.

- Vulgar fractions are not acceptable.

- A comma or similar mark shall not be marked between third and fourth digits of a number.

# 4.	Recommended Numbering Arrangements

## 4.1	First Alternative

Style A (d = 1×10n) where N is an integral multiple of 10d.

Style B (d = 2×10n) where N is an integral multiple of 10d.

Style C (d = 5×10n) where N is an integral multiple of 20d.

## 4.2	Second Alternative

All minor graduation lines (except sub-minor lines) may be numbered incrementally from the preceding principal line, as follows:

(a)	style A – 1d, 2 d, 3 d etc up to 9 d;

(b)	style B – 1d, 2d, 3d, 4 d; and

(c)	style C – 2d, 4d, 6d etc up to 18 d.

## 4.3	Third Alternative

All graduation lines (except sub-minor lines) may be numbered progressively from zero at 1, 5 or 10 graduation intervals.